MesoCam operating instructions

Introduction to the software

Launch Software

- Double click the MesoCam icon on the desktop. The MesoCam user interface is seen in figure 1.

Figure 1: MesoCam user interface.

Identifying Main Features

A: Image Capture

Figure 2: Image capture options.

Live Image: Click live image to display a real-time image of the specimen.

Live Bin: This sets the binning of the live image displayed. The default is set to 1x1; there is also 2x2 and 4x4 available. Higher binning modes (2x2 and 4x4) increase the image capture rate allowing easier focusing/viewing.

Capture Image: Click to begin image capture

Pixel shift: This selects capture mode. The default is set to no pixel shift (Standard(1X1)); there is also the option for 2x2 pixel shift (Hi Res(2x2)) and 3x3 pixel shift (Hi Res(3x3)). Pixel shift imaging increases the spatial resolution of the image by acquiring a grid of sub-images shifted by a fraction of a pixel (1/2 pixel  in 2X2 mode, 1/3 pixel in 3X3 mode and interleaving them into a high resolution image. See notes for resolution details.

Stop: Click to stop the camera. Can be pressed to stop live view or during a capture to cancel.

Exposure time: Can be set between 0.01ms and 7000ms (10µs to 7s). Enter required value and click enter.

Gain: This can be set between 0 and 100. Choose from the dropdown menu. Increasing the gain increases the brightness of the image for a given exposure time. (Note. Increasing the gain also increases image noise so increasing the exposure time is a preferred option for increasing image brightness).

Lens: Selects the objective lens in use. (Note. Only one X4 Mesolens is currently available.)

B: CCD capture area

Figure 3: CCD capture area options.

Full: Click to capture images from the full CCD imaging area,

Use R.O.I.: Click to crop the imaging area to the region of interest defined by the red user drag-able ROI rectangle on the live display.

Set Area: Click the Set Area button to define the imaging region by entering the X and Y dimensions (as a % of the full CCD dimension) into the X Range and  Y Range fields in the Set CCD Area dialog box. Click the OK button select the sub-region.

C: Image mode

Figure 4: Image mode options.

Separate Light Source Images: Feature not in use.

Acquire Z Stack: Allows the user to image a volume. When checked, the sub menu will be displayed, seen in figure 4.         No. sections and Z step size. Number of sections is the number of planes you would like to image. Z step size is the distance between each plane. It can be entered in the pixels text box or µm text box (Recommended input in the µm text box and press enter).

Acquire Time Lapse Series: Allows the user to capture a time lapse. the sub menu will be displayed, seen in figure 4. No. points and At intervals of. No. points is the number of images you want to take. At intervals of is the time delay between each image (input in seconds and press enter).

D: Stage position

Figure 5: Stage position information and control.

Stage position readout: This reads directly from the stage controller and is displayed in µm.

Go To X= : To move the stage to a selected X axis position, enter the position (microns) into the Goto X= field then click the GoTo X= button (or press the Return key)

Go To Y= : To move the stage to a selected Y axis position, enter the position (microns) into the Goto Y= field then click the GoTo Y= button (or press the Return key)

Go To Z= : To move the stage to a selected Z axis position, enter the position (microns) into the Goto Z= field then click the GoTo Z= button (or press the Return key)

E: Contrast

Figure 6b: Contrast sliders tab.

Figure 6a: Display contrast tab.

False colour overlay: Select to change the colour overlay on the image. The default is grey scale. There is also red scale, green scale and blue scale.

Display contrast tab

Full range: The default selection and displays the full range of intensities available for the camera. For the Stemmer camera this is 12 Bit (0-4095).

Best: Click to set the optimal brightness and contrast for the range of pixel intensities within the currently displayed image.

Text box: Dynamic range output. Can be entered manually while focusing the specimen. Type in range in the format of n1 – n2 and press enter. Where n1 is the lower value and n2 is the upper value.

6 x s.d.: Tick the 6 x s.d. option, to make the Best button optimise brightness and contrast to display the range of pixel intensities plus or minus 3 standard deviations of the mean pixel intensity.

Auto adjust: Tick the Auto adjust  option, to optimise optimal brightness and contrast automatically whenever a new image is displayed.

Sliders tab

Contrast: Slide to adjust contrast of displayed image.

Brightness: Slide to adjust brightness of displayed image.

F: Histogram

Figure 7: Pixel intensity histogram and camera readout.

Pixel intensity histogram: intensity distribution of the image displayed in the main window.

Camera readout: Will display if the camera is in live view. When capturing an image the information will be displayed here.

Image Display

H: Mouse hover

Figure 9: Mouse hover information.

Hovering the mouse over the CCD area during live view or when an image is captured will give a readout of pixel information. X: the lateral pixel position, Y: vertical pixel position and I: pixel intensity.

A region of interest (ROI) can be selected by adjusting the bounds of the red rectangle superimposed on  the displayed image .

Left Mouse Button (Zoom In): Click the left mouse button to zoom in on a magnified region of the displayed image centred on the mouse position. To move the zoomed region within image, hold the mouse button down and drag the display.

Right Mouse Button (Zoom Out): Click the right mouse button to zoom out, to revert to displaying a wider region of the image.

Save image

Capture image

Optimise image intensity values

Adjust focus

Using the software

Real-time view

Capture a single image

Real-time view:

- Click live image.

- Ensure the prism is removed from the light path.

Adjust focus:

- Set a short exposure time between 100ms and 500ms.

- Adjust the range of intensities for the displayed image using best or entering a range of intensity values in the text box if required, until the image displayed is visible.

- Use the wheel on the stage control to bring the specimen into focus.

Optimise image intensity values:

- Saturation occurs at 4095 intensity counts. An upper value of 3500 intensity counts is sufficient for imaging with the Stemmer camera.

- Adjust the exposure time, and/or intensity of excitation light to optimise image intensity values. Click best after each adjustment made to update the intensity value text box.

- Once you are satisfied with the intensity values, click full range to ensure you are happy with the optimisation.

Capture image:

- Choose the pixel shift option from the drop down menu (See Resolution details for Stemmer).

- Click capture image. The progress of the image capture will be displayed in the camera readout box.

Save image:

- Click file.

- Click save image to file.

- Choose the location you want to save your file And click save. The file type will be Tiff.

Capture a Z stack

Acquire Z Stack: Tick the Acquire Z Stack option to acquire a stack of images at a series of Z axis planes

Enter the number of sections to be acquired into the Section field and the distance between sections into a Step field. The step size can be entered either as the step size in microns or as multiples of the camera pixel size.

Capture Image: Click Capture Image to capture a stack of high resolution images.

The Z stack begins at the current stage Z position with the stage position incremented by Step microns between images.On completion of the stack the series of images are displayed on screen with each plane selectable for viewing using and Z position scroll bar.

Capture a time lapse series

Tick the Acquire Time Lapse Series option to acquire a series of image at specified time intervals.

Enter the required number of images in the series in the No. Points field and the interval (in seconds) between images in the At Intervals Of field.

Capture Image: Click Capture Image to capture the time series of high resolution images.

On completion of the series, the images are displayed on screen with each plane selectable for viewing using and T position scroll bar.

Saving Images to TIFF files

To save a captured image (or image stack/series) to a TIFF file, select File > Save Image from the main menu to open the Save Image dialog box.

Select a folder to save the files(s), enter a name for the file then click the Save button.

Images are saved in OME-TIFF format.

Saving & Displaying Images in ImageJ

To save a captured image (or image stack/series) to a TIFF file and then open it using Image-J, select File > Save to Image-J from the main menu to open the Save Image dialog box.

Select a folder to save the files(s), enter a name for the file then click the Save button.

The image will be saved to the folder then loaded  into Image-J.

Notes

Resolution details for Vieworks VMC-29MC-5M camera

Values for the Mesolens with the 8x magnification lens attached and single image capture.

Field of View (FOV):  4.4mm x 2.9mm

No pixel shift:

- 6576 x 4384, 28.8 Megapixel.

- 1.486 pixels/µm

- 0.673 µm/pixel

- 0.05 GB

4x pixel shift:

- 13152 x 8768, 115.3 Megapixel.

- 2.973 pixels/µm

- 0.336 µm/pixel

- 0.23  GB

9x pixels shift:

- 19728 x 13152, 259.5 Megapixel.

- 4.425 pixels/µm

- 0.224 µm/pixel

- 0.51 GB